SIPM CONTROLLER DUO

V1.1D

20.07.2011

Serial Communication Interface:

HW Protocol:		 RS485 Half Duplex

Settings:		19200 Baud

8bit

No Parity

1 Stop bit

No Handshake

SW Protocol:

Start Delimiter	Address	  Command      Data            Command	  Data	Stop Delimiter

<	     	      A	        C1	           D1	  	C2	    D2	        >

<	Start Delimiter ASCII Code 0x3C

>	Stop Delimiter ASCII Code 0x3E

A	Address Range 0x00…0xFF

Master	Send	Receive

<			Start Delimiter

AA			Address

AA		Echo If Accepted

CW			Command

CW		Echo If Accepted

D			Data IF Write Command

D		Echo If Accepted

D

D

D

D

CR			Command

CR		Echo If Accepted

D		Data IF Read Command

D

D

> 			Stop Delimiter

Commands:

'<'   Start Delimiter	ASCII Code 0x3C

'>'   Stop Delimiter ASCII Code 0x3E

Address hex value 0x00…0xFF

'I'   Read Temperature

Return example	I22.1

*

'A’   Set Bias Voltage @ 25 deg C Sensor A (0000...3FFF  5 mV/count)

'B’   Set Bias Voltage @ 25 deg C Sensor B (000...3FFF  5 mV/count)

Sets the reference bias voltage @ 25 deg C or the absolute bias voltage if the progressions coefficient is set to zero. Range 0x0000…0x3FFF  5mV/count

'C'   Set Bias Voltage Progressions Coefficient 00...5F  mV/K

'D'   Read Temperature Adjusted Voltage (Calculated) Sensor A

'E'   Read Temperature Adjusted Voltage (Calculated) Sensor B

'F'   Read Progression Coefficient

Return example 	D2

*

'I'   Read Temperature

'H'   Print This List

Specifications:

Voltage Regulator: 			two channels 0…80V  5mV step

Temperature Measure Range:	 -30,0 …+45,0  ˚C resolution 0,1 K

Power consumption:		46mA@3,3V; 36mA@-3,3V; 0,5mA@80V

60mA@5V; 40mA@-5V; 0,5mA@80V

Kalibration:

1. Das erste Zeichen nach dem einschalten und der akzeptierten Adresse muss ein ‘Z’ sein um in den             Kalibrations Modus zu gelangen ( < ‘address‘‘Z‘………………………>)

2. Mit ‘D‘1F4 Sollspannung 10,00V vorgeben, mit ‘Y‘ und ‘+‘ oder ‘-‘ Ist Spannung auf 10,00V einstellen, mit ‘S‘ Offset Einstellung verlassen.

3. Mit ‘D’DAC Sollspannung auf 70,00V vorgeben, mit ‘X‘ und ‘+‘ oder ‘-‘ Ist Spannung auf 70,00V einstellen, mit ‘S‘ Verstärkungs Einstellung verlassen.

4. Schritt 2. Und 3. Wiederholen.